# TYPE APPROVAL CHECKLIST

## Digital Non-Automatic Weighing Instruments (NAWI)

### Modular Assessment: Load Cell + Indicator + Legally Relevant Software

# SECTION 1 — ADMINISTRATIVE INFORMATION

| Item | Requirement | Evidence Submitted | Complies (Y/N) | Remarks |
| --- | --- | --- | --- | --- |
| 1.1 | Applicant name & contact details |  |  |  |
| 1.2 | Manufacturer name & address |  |  |  |
| 1.3 | Instrument designation / model |  |  |  |
| 1.4 | Accuracy class (I, II, III, IIII) |  |  |  |
| 1.5 | Max, Min, e, d declared |  |  |  |
| 1.6 | Multi-range / multi-interval declared |  |  |  |
| 1.7 | Technical file complete |  |  |  |

# SECTION 2 — LOAD MEASURING DEVICE (Load Cells)

### Reference: SANS 1649 Annex F.2

| Ref | Requirement | Evidence | Y/N | Remarks |
| --- | --- | --- | --- | --- |
| F.2.1 | Load cell has valid certificate of conformity (R60) | Certificate |  |  |
| F.2.1 | Only SH or CH tested load cells used | Certificate class |  |  |
| F.2.2 | Accuracy class compatible with instrument class | Calculation |  |  |
| F.2.3 | pLC within 0.3–0.8 (default 0.7 if not declared) | Technical file |  |  |
| F.2.4 | Temperature limits declared (or −10 °C to 40 °C default) | Certificate |  |  |
| F.2.5 | Emax ≥ Q × Max × R/N | Calculation sheet |  |  |
| F.2.6 | Emin ≤ DL × R/N | Calculation sheet |  |  |
| F.2.7 | nLC ≥ n (or ni for multi-interval) | Calculation |  |  |
| F.2.8 | DR requirements satisfied (if applicable) | Technical data |  |  |
| F.2.9 | vmin ≤ e1 × R/√N | Calculation |  |  |
| F.2.10 | Input resistance compatible with indicator (RLmin/RLmax) | Compatibility test |  |  |

# SECTION 3 — INDICATING DEVICE

### Reference: SANS 1649 Clause 4.4

| Ref | Requirement | Evidence | Y/N | Remarks |
| --- | --- | --- | --- | --- |
| 4.4.1 | Change of indication ≤ 1 s | Test report |  |  |
| 4.4.2 | Stable equilibrium logic defined | Manufacturer doc |  |  |
| 4.4.2 | Printing inhibited when unstable | Test |  |  |
| 4.4.2 | Storage inhibited when unstable | Test |  |  |
| 4.4.3 | Extended indication rules complied | Functional test |  |  |
| 4.4.4 | Additional indications not ambiguous | Demo |  |  |
| 4.4.5 | Printing legible (≥ 2 mm), unit properly shown | Print sample |  |  |
| 4.4.6 | Memory storage protected | Test |  |  |

# SECTION 4 — SEPARATELY TESTED INDICATOR (If Modular)

### Reference: SANS 1649 Annex F.3

| Ref | Requirement | Evidence | Y/N | Remarks |
| --- | --- | --- | --- | --- |
| F.3.1 | Valid indicator certificate | Certificate |  |  |
| F.3.2 | Class compatible with instrument | Review |  |  |
| F.3.3 | pind within 0.3–0.8 (default 0.5) | Technical file |  |  |
| F.3.4 | Temperature limits declared | Cert |  |  |
| F.3.5 | nind ≥ n (or ni) | Calculation |  |  |
| F.3.6 | Multi-interval/multiple range functions certified | Cert |  |  |
| F.3.7 | Electrical compatibility (∆umin, RLmin, RLmax) | Tech data |  |  |

# SECTION 5 — LEGALLY RELEVANT SOFTWARE

### Reference: SANS 1649 Clause 5.5 + OIML D 31

## 5.1 Software Identification

| Ref | Requirement | Evidence | Y/N | Remarks |
| --- | --- | --- | --- | --- |
| 5.5.1 | Software declared embedded or loadable | Declaration |  |  |
| 5.5.1 | Legally relevant functions described | Technical doc |  |  |
| 5.5.1 | Software version/checksum available via display | Demo |  |  |
| D31 5.x | Software ID linked to approval documentation | Review |  |  |

## 5.2 Protection Against Modification

| Ref | Requirement | Evidence | Y/N | Remarks |
| --- | --- | --- | --- | --- |
| 5.5.2 | Software protected against intentional change | Security design |  |  |
| 4.1.2.4.4 | Evidence of intervention (event counter or equivalent) | Demo |  |  |
| D31 | Audit trail non-erasable | Test |  |  |
| D31 | Parameter changes logged (old/new value + time) | Demo |  |  |
| D31 | Audit log printable/displayable | Test |  |  |

## 5.3 Protective Interface & Separation

| Ref | Requirement | Evidence | Y/N | Remarks |
| --- | --- | --- | --- | --- |
| 5.5.2.2.3 | Protective interface implemented | Interface spec |  |  |
| D31 6.3.2 | Legally relevant modules clearly defined | Architecture doc |  |  |
| D31 | Non-legal software cannot influence results | Test/demo |  |  |
| D31 | Communication commands limited & documented | Interface table |  |  |

## 5.4 Shared Indications

| Ref | Requirement | Evidence | Y/N | Remarks |
| --- | --- | --- | --- | --- |
| D31 | Legal data always clearly distinguishable | Demo |  |  |
| 4.4.4 | Non-weighing values clearly identified | Demo |  |  |

# SECTION 6 — MARKING & IDENTIFICATION

| Requirement | Evidence | Y/N | Remarks |
| --- | --- | --- | --- |
| Data plate present (manufacturer, class, Max, Min, e) | Photo |  |  |
| Serial number unique | Inspection |  |  |
| Software ID accessible | Demo |  |  |
| Verification marks/sealing provisions provided | Inspection |  |  |

# SECTION 7 — OVERALL CONFORMITY DECISION

| Item | Outcome |
| --- | --- |
| All mandatory clauses satisfied | ☐ Yes ☐ No |
| Type Approval Recommendation | ☐ Approve ☐ Reject ☐ Further info required |
| Conditions of approval (if any) |  |
| Assessor name |  |
| Date |  |

# OPTIONAL ADD-ON — CALCULATION VERIFICATION BLOCK

You may include a dedicated compatibility calculation sheet section for:

- pLC + pind allocation

- nLC + nind compatibility

- vmin compliance

- Emax / Emin formula compliance

- RL compatibility